Switch Access Pod User Guide

NOTE: Please read these instructions before connecting this product up to your games console.

These instructions have been prepared in order that the product is not used outside of its limitations, and to ensure the user obtains maximum satisfaction from its use.

# CONTENTS

FRONT SHEET	PAGE 1

CONTENTS PAGE	PAGE 2

1	INTRODUCTION	PAGE 3

2	LIMITATIONS OF USE	PAGE 3

3	WARNINGS	PAGE 3

4	DESCRIPTION OF UNIT	PAGE 4

5	OPERATING INSTRUCTIONS 	PAGE 5
6	PROBLEM SOLVING	PAGE 5

7	CLEANING INSTRUCTIONS	PAGE 5
8	TECHNICAL SPECIFICATION 	PAGE 6

9	ENHANCED GAMING 	PAGE 7
10	Declaration & LEPMIS Contact details	PAGE 8

1	INTRODUCTION

This manual provides the necessary information for the installation and use of the LEPMIS PS3-SAP Switch Access Pod. Please read this manual before attempting to connect and power up the PS3-SAP with your games console. The manual also provides basic trouble-shooting information. If you require further help, please contact LEPMIS with fault details and include unit Type, Date of Manufacture and Serial Number as described on the labelling of the unit.

PS3-SAP is classed as an “accessory to a Medical Device” for the purposes of the European Medical Devices Directive 93/42/EEC 14th June 1993, as it compensates for an injury or handicap. Its purpose is to enable gamers of all physical ability to play games on the SONY™ PS3 Playstation (or Microsoft™ Xbox 360 via an additional adapter) where a standard controller would be a barrier to access. Any user switch and user analogue joystick can be used for control via the unit. An adapter which corresponds to the PS3-SAP joystick analogue port specification (in this manual) will be needed to connect your analogue joystick. LEPMIS or a suitably qualified technician can produce the adaptor, or LEPMIS can provide a suitable joystick from our wide range of uniquely sensitive and responsive types.

LEPMIS is able to perform a clinical assessment for and manufacture suitable switch set-ups and joystick units to meet your particular needs. If in any doubt, contact LEPMIS by one of the methods at the end of this document.

- LIMITATIONS OF USE

The unit is reasonably well protected from the effects of dust, water and EMI, but there are no particular features to prevent water entering the unit. Do not soak the unit or attempt to “wash” it in any way. Wipe sparingly with a barely damp cloth and mild detergent if really necessary.	

The unit is intended to be used in a home or school environment, and due to its lack of full environmental protection, must not be subjected to exposure to the elements. The unit is rated for continuous operation but you are advised to consult the games console documentation on the length of time the unit should be used or left ON.

Operating temperature range, as per your game console operating instructions.

# WARNINGS

# No unauthorised access or dismantling of the unit allowed! Tampering will terminate your warranty. Compentent technicians requiring internal access to PS3-SAP must apply to the manufacturers – LEPMIS, in writing stating their aim and seeking authorisation to proceed.

# It is not recommended that you play computer games for extended periods of time without breaks, as gaming can become an addictive habit, leading to emotional and physical changes to the player.	

USE YOUR GAMES CONSOLE SENSIBLY !

4	DESCRIPTION OF UNIT

PS3-SAP replicates all functions of the SONY PS3 handset through seventeen 3.5mm stereo jack sockets which accept user switches, and two 9 way “D” type sockets for the users joysticks. A USB socket accepts a USB cable from the games console for power and to pass the users control to the games console. To help setting up the game, the PS button is replicated on the top of the box, and the Start and Select buttons can be found on the back of the box next to their switch sockets. On the underside of the unit is a small hole for the RESET function.	

On the front face of the unit are the sockets for Up, Down, Left, Right direction button inputs, and the Triangle, Cross, Box, Circle symbol button inputs.

At the left end of the unit are the Left and Right Joystick analogue ports. NOTE these do not use the same pin configuration as some common 9 way “D” type game joysticks (see the technical information on PS3-SAP analogue port specifications in this manual). Also on the left of the box are sockets for the L3, R3, and L2 shoulder button inputs.

On the back of the unit are the sockets for L1, R1 shoulder button inputs, plus Start, Select and PS.

On the right hand side of the unit is the socket for R2 and the games console USB cable.

LEPMIS recommends the use of STEREO 3.5mm jack plugs on the user’s switches. MONO jack plugs are perfectly acceptable and will almost certainly be what you find attached to your commercially available switches. Stereo jack plugs however, when wired up as detailed later, will preserve the electrical separation of the outer metal ring of the sockets from the electronic circuits inside. This is a reliability feature and helps to protect the electronic circuits from the effects of ESD, (Electro-Static Discharge), which is the “shock” you sometimes get when walking across a carpet and then touching a metal or conductive surface.

5	OPERATING INSTRUCTIONS

Plug the USB cable from the games console into the socket on the right hand side of the unit. Note: the unit will only work so long as the USB cable is plugged in, as the unit does not have an internal battery.

The Start, Select and PS buttons and switch sockets work in the same way as on the handset.

Switches can be plugged into or out of the unit at any time. Please note however that the action of plugging or unplugging will cause a momentary switch closure, interacting with any game you may be playing. Pressing Up / Down together will result on NO switch action. Similarly with Left / Right.

An analogue joystick can also be inserted or removed at any time, but you may see some slight joystick interaction with the game until all of the pins of the plug are fully mated or fully removed.

LEPMIS is happy to advise on and supply suitable switches and joysticks. Additionally, we can assess and manufacture a custom configuration to exactly meet a difficult user requirement. Assessment may be dependent on the location of the user.	


This symbol       	means that the PS3-SAP must be re-cycled when no longer required. DO NOT throw away in the general refuse. Please request from LEPMIS a pre-paid envelope in which to return it for appropriate re-cycling.

6	PROBLEM SOLVING

Always check that your switches or joystick are working correctly before suspecting this unit, as the wires on many commercial switches are known problem areas.

A wide range of suitable switches is available from www.inclusive.co.uk. LEPMIS recommend you specify a switch designed for repeated use, that is waterproof, that has a reinforced cable, and (ideally) has a stereo jack plug. If in doubt, we can help to choose the right switch for you.	

The PS3-SAP sockets are electronically resilient, and even with repeated connection and disconnection of the plugs should last many years of heavy use before they need replacing. AVOID getting fluids or foreign material into the sockets as they will either corrode or become jammed up, and stop working. Failure of the sockets in this way is not covered by the warranty. If you suspect PS3-SAP sockets are working intermittently or analogue ‘drift’ (game character moves without user input) then please contact LEPMIS (email, phone, address are listed on the last page of this manual) for repairs or maintenance.

7 	CLEANING INSTRUCTIONS

When necessary, the unit should be cleaned with a barely damp cloth, using only a tiny amount of household detergent. DO NOT USE abrasive products or solvents.

# 8	TECHNICAL SPECIFICATION

Size:		18cm wide,  by 12.7cm deep, by 4cm high

Weight:		0.375 kg

Operating Voltage:	5v dc, (from the games console)

# User Switch Jack plug:	3.5mm Mono or Stereo (preferred), normally open

# Maximum recommended switch contact resistance – 100 ohms. 
Absolute maximum contact resistance, 390 ohms.

NOTE: 	The electrical connections for the user switches and joysticks are isolated from the electronic circuits within the PS3-SAP and the games console.

# Stereo Plug:

# Connect your switch between the “TIP” and “RING” terminals of the connector.
Connect the screen of your cable to the BODY contact, ( if cable has a screen).

# NOTE: 	0v will be connected to the RING and the +Ve feed will be connected to the TIP. The action of the switch is to bring the TIP down to 0v. +Ve feed output impedance 2k7 ohms.

# Mono Plug:

# Connect your switch between the “TIP” and “BODY” terminals of the connector.

# NOTE: 	0v will be connected to the BODY of the plug when the plug is mated. As before +Ve feed will be connected to the TIP and the action of the switch is to bring the TIP down to 0v. All switch inputs have a common 0v connection.

# Analogue Port Specification	(Nominally 10kΩ, middle position = no movement)

# Up-Down track

# 1	Up

# 2	Down

# 6	Wiper

# Left-Right track

# 3	Left

# 4	Right

# 8	Wiper

# 7	L3/R3 button input

# 5	+5v @ 400mA max

# 9	0v (L3/R3 button common)

# Output impedance, pins 1,2,3,4 - 270 ohms. Input impedance, pins 6,8 -  1k ohms

# Port pins are protected from short circuit and ESD.

# 9	ENHANCED GAMING

# Through PS3-SAP you can add many different connection options to help with access to gaming. For example, a simple switched analogue joystick can be implemented as follows:

# You can also use a slide potentiometer so that the game player has the ability to set and hold an analogue joystick in a particular position – as a throttle in a car race game, say.

# More than one switch function socket can be connected together through a multi-pole switch, so that those awkward switch combinations can be achieved by pressing just one switch.

# An auto repeat feature can be implemented with either a motor driven mechanical relay or an electronic timer circuit. This can allow the user to, say, continuously fire automatically whilst concentrating on guiding the character around the screen.

# Using a simple “splitter” more than one switch can be connected to the same switch socket or function so that more than one person can access that game function.

# LEPMIS have a wide range of accessories to help gamers of all abilities get maximum enjoyment through PS3-SAP. We are happy to offer advice or manufacturing services on an individual basis.

# 10	Declaration & LEPMIS Contact details

This User Guide has been written for the:-

LEPMIS PS3-SAP Switch Access Pod.

# which is designed to meet the general requirements of:-

the Medical Devices Directive 93/42/EEC  14 June 1993.

and conforms to the requirements of BSEN 60529 to a level of IP30

Supplied by:

# LEPMIS Game Access

# 24 Laurel Crescent

# Long Eaton

# Nottingham

# NG10 3NN

# Tel:				0771  495 8061 (Geoff Harbach)

# e-mail:				Support@LEPMIS.co.uk

# web: 				www.lepmis.co.uk

# We hope you enjoy using this product, but please do so safely!

# End of Document

# Text boxes

|	|	|	|

8

4

2

Right

UP	DOWN	LEFT	RIGHT

|	|	|	|

6

1

Up

Left

LEPMIS Game Access Solutions

Unique Solutions for Unique People

START	SELECT	PS	L1	R1

|	|	|	|	|	|	|

3

Down

LEFT JOYSTICK	L3	RIGHT JOYSTICK	R3	L2

|	|	|	|	|

R2	USB

|

Body

Ring

Tip

3.5mm Stereo Jack Plug

3.5mm Mono Jack Plug

Tip

1

Body

2

3

4

5

9

8

7

6

4    (2)

8    (6)

3    (1)